Appendix Four

ANTHROPOMETRIC TECHNIQUES

| The UNICEF Electronic Scale 890 |  |
| --- | --- |

The UNICEF Electronic Scale

The UNICEF Electronic Scale was designed to assist health workers monitor the weight of children and pregnant women.

The scale makes weighing fast, easy, and accurate.  It can be used in two ways:

- Pregnant mothers or older children can line up for weighing, stepping on the scale one after the other.

- Babies and very small children can be weighed while being held in the arms of a mother or helper.  This second method of weighing is called “tared weighing.”

The scale is powered by long-lasting lithium batteries.  These will complete at least one million weighing cycles, or 400 weighings every day of the year (except weekends and holidays) for at least ten years.  The batteries and the electronic “heart” of the scale are in a sealed unit to withstand damage from heart, humidity, and dust. The solar cell is used only to turn the scale on and to tare the scale.

The scale switches off automatically if it is not used for two minutes. This helps increase battery life.

## Preparing the UNICEF Electronic Scale for use

- Place the scale on a hard, level surface (wood, concrete, or firm earth).  Soft or uneven surfaces may cause small errors in weighing.

- The scale will not function correctly if it becomes too warm.  It is best to use the scale in the shade, or indoors.  If the scale becomes hot and does not work correctly, place it in a cooler area and wait 15 minutes before using again.

- The scale must adjust to changes in temperature.  If the scale is moved to a new site with a different temperature, wait for 15 minutes before using the scale again.

- Handle the scale carefully:

- Do not drop or bump the scale.

- Do not weigh loads with a total weight of more than 150 kg.

- Do not store the scale in direct sunlight or other hot places.

- Protect the scale against excess humidity or wetness.

- Do not use the scale at tempera-tures before 0ºC or above 45º C.

### Cleaning

To clean the scale, wipe surfaces with a damp cloth. Never put the scale into water.

### Storage

Do not store the scale in direct sunlight or other hot places.

Weighing an Infant or Young Child Held by a Health Worker or Helper (Tared Weighing)

Note:

is used to symbolize the mother and child image that appears in the left side of the display whenever the scale has been tared.

- Turn the scale on by covering the solar cells for less than one second.

The display should show  first, then .

Wait until the display shows   before stepping on the scale.

Cover the solar cells for less than one second.

| The scale is preparing itself for use. | The scale is ready for use. |
| --- | --- |

- Ask your helper to stand on the scale.  Make sure the solar cells are not covered by a skirt or by the person’s feet.  Your helper’s weight will appear in the display within two seconds.

NOTE:

The person being weighed must stand still on the scale.

|  | Your helper’s weight appears in the display. |
| --- | --- |

- With your helper standing still on the scale, cover the solar cell for less than one second.  The display will read  . The small picture of the mother and baby means the scale is adjusting (taring) itself to ignore the helper’s weight and preparing itself to weigh only the baby.

Cover the solar cells for less than one second to tare the helper’s weight.  Once tared, the display will show .

- The helper can now get off the scale to get the baby or the mother can hand the baby to her.

If the helper gets off the scale to get the baby, the display will show .

If the helper gets off the scale, the display shows that the scale has been tared for the helper’s weight.

- After the helper steps back onto the scale and holds the baby, only the weight of the baby will be displayed.

Record the baby’s weight.

NOTE:

The scale will continue to show only the weight of the baby until the solar cell is covered or the helper returns the baby to its mother.

Now the helper can hold the baby and get back on the scale.  Only the baby’s weight will show in the display.

- After the baby is given back to its mother, the display will again show      (as long as the helper remains standing on the scale).  If the helper steps off the scale to get another baby, the display will show .

If the helper gives the baby to someone else to hold, the tared scale will again display .

- Repeat steps 4 and 5 to weigh another baby.

Remember:  The scale switches off automatically two minutes after the last weighing.  If this happens, follow the instructions to turn it on again.

### Taring Weights:
Important Points

- The weight of the person who will hold the child must be displayed (and then tared) before the child is given to them for weighing.

- The same person whose weight is tared must also hold the child.

- The weight of the child being weighed must be more than 2 kg if the helper stays on the scale to receive the child.

- If the helper gets off the scale to get a child while the display shows   (tared weighing mode), the next child can weigh less than 2 kg.

- The tare can be deactivated by quickly covering the solar cell or by waiting until the scale switches off automatically.

- The scale will always display  if the new load weighs less than the tared weight.

- If there is too much movement on the scale during measurement, the display will switch between  and  until the load becomes stable.

Possible reasons for the scale not taring weight:

- There was no weight on the scale to tare. Put some weight on the scale and try again.

- The solar cell was not covered completely.

- The solar cell was covered for more than one second.  Try covering it for less than one second.

- It is too dark.  Put the scale in a brighter place.

- The load weighs more than 120 kg.  Use a lighter load.

What to Do if the Scale Display Shows . . .

E01:

The scale has to adjust itself.  Get off the scale and wait until E01 no longer appears.

E02
and switches off automatically:

Be sure there is no load on the scale and try to start the scale.

E03
and switches off automatically:

The scale is either too cold or too hot. Move it to a different place with the temperature between 0ºC and 45ºC.  Wait 15 minutes for it to adjust to the new temperature, then start the scale.

E04
after measuring:

The load is too heavy (more than 150 kg.).  Get off the scale and reduce the load.

E05
for a few seconds after trying to start the tare function:

The load is too heavy for taring (more than 120 kg.).  Get off the scale and reduce the load.

##### Instructions for Supervisors

I. 	Procedures and Precautions Before Measuring

A.	Layout of the Procedures

Each step of the measurement procedures is directed at specific participants, who are named in bold letters at the beginning of each step: e.g., “Measurer”,  “Assistant”, etc.

B.	Two Trained People Required

Two trained people are required to measure a child’s height and length.  The measurer holds the child and takes the measurements.  The assistant helps hold the child and records the measurements on the questionnaire.  If there is an untrained assistant such as the mother, then the trained measurer should also record the measurements on the questionnaire.  One person alone can take the weight of a child and record the results if an assistant is not available.

C.	Measuring Board and Scale Placement

Begin to observe possible places where the electronic scale and board can be positioned as soon as you walk towards a sample household.  Be selective about where you place the measuring board and electronic scale.  It is best to measure outdoors during daylight hours.  If it is cold, raining or if too many people congregate and interfere with the measurements, it may be more comfortable to weigh and measure a child indoors.  Make sure there is adequate light.

D.	Age Assessment

Before you measure, determine the child’s age.  If the child is less than two years, measure length.  If the child is two years of age or older, measure height.   If accurate age is not possible to obtain, measure length if the child is less than 85cm.  Measure height if the child is equal to or greater than 85cm.

E. 	When to Weigh and Measure

Weigh and measure after verbal information has been recorded on the questionnaire.  This will allow you to become familiar with the members of the household.  DO NOT weigh and measure at the beginning of the interview, i.e. as soon as you enter a household, which would be more of an upsetting intrusion.

F.	Weigh and Measure One Child at a Time

If there is more than one eligible child in the household, complete the entire questionnaire, including the weighing and measuring of one child.  Then proceed with the next eligible child’s questionnaire in the household.  DO NOT weigh and measure all the  children together.  This can easily cause confusion and will create a greater chance for error such as recording one child’s measurements on another child’s questionnaire.  Return measuring equipment to their storage bags immediately after you complete the measurements for each household.

G.	Control the Child

When you weigh and measure, you must control the child.  The strength and mobility of even very young children should not be underestimated.  Be firm yet gentle with children.  Your own sense of calm and self confidence will be felt by the mother and the child.

When a child has contact with any measuring equipment, i.e. on a measuring board, on the electronic scale, you must hold and control the child so the child will not trip or fall.  Never leave a child alone with a piece of equipment.  Always have physical contact with the child except when you must let go of a child for a few seconds while taking the weight.

H. 	Coping With Stress

Since weighing and measuring requires touching and handling children, normal stress levels for this type of survey work are higher than for surveys where only verbal information is collected.

Explain the weighing and measuring procedures to the mother, and to a limited extent, the child, to help minimize possible resistance, fears or discomfort they may feel.   You must determine if the child or mother  is under so much stress that the weighing and measuring must stop.  Remember, young children are often uncooperative; they tend to cry, scream, kick and sometimes bite .  If a child is under severe stress and is crying excessively, try to calm the child or return the child to the mother for a moment before proceeding with the weighing and measuring.

Do not weigh or measure a child if:

- The mother refuses.

- The child is too sick or too distressed.

- The child is physically deformed which will interfere with or give an incorrect measurement.  To be kind, you may want to measure such a child and make note of the deformity on the questionnaire.

I. 	Recording Measurements and Being Careful

Record the measurements in pen.  If you make an error, completely erase the error and rewrite the correct numbers.  Keep objects out of your hands and pens out of your mouth, hair or breast pocket when you weigh and measure so that neither the child nor you will get hurt due to carelessness.  When you are not using a pen, place it in your equipment pack, pen case or on the survey form.  Make sure you do not have long fingernails.  Remove interfering rings and watches before you weigh and measure.  Do not smoke when you are in a household or when you weigh and measure.

J. 	Strive for Improvement

You can be an expert measurer if you strive for improvement and follow every step  of every procedure the same way every time.  The quality and speed of your measurement will improve with practice.  You may be working with a partner to form a team.  If so, you will be responsible for not only your own work, but also for the quality of work of your team.

You will be required to weigh and measure many children..  Do not take these procedures for granted even though they may seem simple and repetitious.  It is easy to make errors when you are not careful.  Do not omit any steps.  Concentrate on what you are doing.

Nutritional Status Measurement

Summary Procedures

A.	Child Height Summary Procedure (Illustration 1)

1.	Measurer or Assistant:   Place the measuring board on a hard flat surface against a wall, table, tree, staircase, etc.  Make sure the board is stable.

2.	Measurer or Assistant:  Ask the mother to remove the child’s shoes and unbraid any hair that would interfere with the height measurement.  Ask her to walk the child to the board and to kneel in front of the child (if she is not the assistant)

3. 	Assistant:  Place the questionnaire and pencil on the ground (Arrow 1)  Kneel with both knees on the right side of the child (Arrow 2).

4.	Measurer:   Kneel on your right knee only, for maximum  mobility, on the child’s left side  (Arrow 3).

5.	Assistant:   Place the child’s feet flat and together in the centre of and against the back and base of the board.  Place you right hand just above the child’s ankles on the shins (Arrow 4), your left hand on the child’s knees (Arrow 5) and push against the board.  Make sure the child’s legs are straight and the heels and calves are against the board (Arrows 6 and 7).  Tell the measurer when you have completed positioning the feet and legs.

6.	Measurer:   Tell the child to look straight ahead at the mother if she is in front of the child.  Make sure the child’s line of sight is level with the ground (Arrow 8).  Place your open left hand on the child’s chin.  Gradually close your hand (Arrow 9).  Do not cover the child’s mouth or ears.  Make sure the shoulders are level (Arrow 10), the hands are at the child’s side (Arrow 11), and the head, shoulder blades and buttocks are against the board (Arrows 12,13 and 14).  With your right hand, lower the headpiece on top of the child’s head.  Make sure you push through the child’s hair (Arrow 15).

7.	Measurer and Assistant:  Check the child’s position (Arrow 1-15).  Repeat any steps necessary.

8.	Measurer:   When the child’s position is correct, read and call out the measurement to the nearest 0.1cm.  Remove the headpiece from the child’s head, your left hand from the child’s chin and support the child during the recording.

9.	Assistant:   Immediately record the measurement and show it to the measurer.

NOTE:  If the assistant is untrained, the measurer records the height.

10.	Measurer:   Check the recorded measurement on the questionnaire for accuracy and legibility.  Instruct the assistant to erase and correct any errors.

#### Illustration 1

B. 	Child Length Summary Procedure (Illustration 2)

1.	Measurer or Assistant:  Place the measuring board on a hard flat surface, i.e. ground, floor or steady table.

2.	Assistant:   Place the questionnaire and pencil on the ground, floor or table (Arrow 1).  Kneel with both knees behind the base of the board, if it is on the ground floor. (Arrow  2).

3.	Measurer:   Kneel on the right side of the child so that you can hold the footpiece with your right hand (Arrow 3).

4.	Measurer and Assistant:   With the mother’s help, lay the child on the board by doing the following:

Assistant:   Support the back of the child’s head with your hands and gradually lower the child onto the board.

Measurer:   Support the child at the trunk of the body.

5.	Measurer or Assistant:   If she is not the assistant, ask the mother to kneel on the opposite side of the board facing the measurer to help keep the child calm.

6.	Assistant:   Cup your hands over the child’s ears (Arrow 4).  With your arms comfortably straight (Arrow 5), place the child’s head against the base of the board so that the child is looking straight up.  The child’s line of sight should be perpendicular to the ground (Arrow 6).  Your head should be straight over the child’s head.  Look directly into the child’s eyes.

7.	Measurer:   Make sure the child is lying flat and in the centre of the board (Arrows 7).  Place your  left hand on the child’s shins (above the ankles) or on the knees (Arrow 8).  Press them firmly against the board.  With your right hand, place the footpiece firmly against the child’s heels (Arrow 9).

8.	Measurer and Assistant:   Check the child’s position (arrows 1-9).  Repeat any steps as necessary.

9.	Measurer:   When the child’s position is correct, read and call out the measurement to the nearest 0.1 cm.  Remove the footpiece, release your left hand from the child’s shins or knees and support the child during the recording.

10.	Assistant:   Immediately release the child’s head, record the measurement, and show it to the  measurer.

NOTE:   If the assistant is untrained, the measurer records the length on the questionnaire.

11.	Measurer:   Check the recorded measurement on the questionnaire for accuracy and legibility.  Instruct the assistant to erase and correct any errors.

#### Illustration 2

# Footnotes

If the assistant is untrained, e.g. the mother, then the measurer should help the assistant with the height procedure.

If the assistant is untrained, e.g. the mother, then the measurer should help the assistant with the length procedure.